Comparative study to understand the clinically efficiency of Vigo Smart Heart (VSH) against traditional holter in Indian Arrhythmia patients

Katika Madhumohan1*, Velugonda Nagaraju2, Bandaru Rajiv3, Maddur Srinivas1, Agrawal Sony4, Yerradoddi Sushma4, Lavudi HariKrishna Naik5, Jakinala Parameshwar5

1CRL, ESIC Medical College and Hospital Hyderabad, Sanath Nagar, Hyderabad, Telangana - 500038

2Nizam's Institute of Medical Sciences, Molecular Diagnosis, Punjagutta, Hyderabad, Telangana-500082

3Department of Medicine, SIC Medical College and Hospital Hyderabad, Sanath Nagar, Hyderabad, Telangana – 500038

4Department of Cardiology, SIC Medical College and Hospital Hyderabad, Sanath Nagar, Hyderabad, Telangana – 500038

5MVRDL, ESIC Medical College and Hospital Hyderabad, Sanath Nagar, Hyderabad, Telangana - 500038

*Corresponding Author: Dr. Madhumohan Katika

E-mail: drmadhumohannizams@gmail.com

Telephone number: +919182015668

Abstract

Arrhythmia is an irregular hertbeat which leads to severe  heart complications and it is the most common action of Cardio Vascular Diseases. India represents 31% of global deaths and need for accurate diagnosis and monitoring of Arrhythmia has not been addressed. In view, we developed a novel device Vigo Holter (VSH) that is connected to cloud and IoT based platform designed as an easy wearable for the patient. It records continuous ECG and HR to predict the changes in the heart. Compared a Vigo Holter against the traditional holter monitoring in 51 volunteers for 24 hrs with asymptomatic and symptomatic subjects. We evaluated patient compliance, analyzable signal time interval to arrhythmia detection, and diagnostic yield.In total 51 participants we found 46 reports with equivalent result where as the conditions (Second Degree Mobitz Type I block, First Degree AV block, IVCD and SVT episodes) identified by VSH  . Importantly, ECG wave quality in reports with differences is same in both recordings and the total diagnostic yield was 39%. Total Noise in Traditional Holter was 1301 minutes whereas in Vigo Holter was 990 minutes. Total Analyzable time in Vigo Holter was 99.3% whereas in Traditional Holter was 90.22%. VSH reports clearly explained that no lead detachments and noise resultant from the wire entanglements leading to low noise and highly analyzable time. We demonstrated that VSH is very much needed and useful for people and doctors to detect arrythmia with highest accuracy and to avoid physical interaction with the patient during COVID-19 pandemic.

Keywords: Arrhythmia; Cardio Vascular Diseases; ECG; Traditional holter; Vigo Smartheart

1 INTRODUCTION

Arrhythmia [1,2]  is the terminology used to indicate frontline situation of abnormal heart beat if untreated that can lead to irreparable damage or even death and it is the most common manifestation of Cardio Vascular Diseases (CVD). Life style, food habits[3] and individual genetics are root causes of CVDs. The total  number of CVD are estimated to be 54.6 Million/year and death rates are 2 million/year in India. In 2016, 17.9 million CVD deaths were recorded in india that represented 31% of global deaths. Having set the CVD burden bars high, limited research has been done to assess the prevalence and prognosis of arrhythmias[2] in India. Therefore, it is imperative to record arrhythmias in healthy people vs. comorbid patient subjects to design tailor therapies for CVD.

The development chronology of arrhythmia diagnostic tools are  ECG, Traditional Holter, Event monitors, Loop recorders[2], Mobile cardiac Telemetry and continious ECG monitoring solutions. ECG is the most common technology used to diagnose cardiac arrhythmia due to its low cost and maintenance. Traditional Holter monitor for ECG monitoring was developed by Norman J. Holter and Bill Glasscock in 1949. ECG combined with technology of deep learning methodologies[4] has been a boon for automatic detection of arrhythmias [7]. The next level deep learning improvement trials included for up to date deep neural networks along with handcrafted features of ECG to arrive at multi resolution representation (MRR) of ECG signals[5] [8]. Another group performed a five class ECG dataset containing 1,00,022 heart beats to analyze with advanced deep learning techniques and presented good correlating results [9].

Recently, a personalized arrhythmia monitoring platform[4] was developed using novel method called discrete orthogonal stockwell transform (DOST) to capture significant time-frequency coefficient that constitute a feature set representing each ECG signal [10]. In the recent decade (2010 – 2020) a systematic review of 191 publications related to deep learning based ECG analysis was performed and found that hybrid architecture of convulutional neural network and recurrent neural network ensemble using expert features yields the best results [11]. Another systemic literature search in MEDLINE and EMBASE revealed that Holter monitors[2] have less diagnostic yield (33 to 35%) than event recorders; recorders with automatic triggering detect more arrhythmias (72 to 80%) than patient-triggered devices (17 to 75%).

Implantable devices are used for prolonged monitoring periods in patients with infrequent symptoms or unexplained syncope. Latest innovations redefined the ECG monitoring by making the devices compact, minimalistic and convenient for the patients. These innovative ECG monitoring solutions need to establish their clinical accuracy and efficiency [6,7] in varied conditions especially in Indian settings. However, challenges and problems still exist that are related to interpretability, scalability, and efficiency should be addressed. Although it is needless to mention that choice of device depends on specific symptoms of the patient, it is important to have a device with added features capable of performing unconventional detection of arrhythmias compared to traditional holters.

Thus, in this study we intended to compare a remote continious ECG monitoring against the traditional holter monitoring. We aim to evaluate patient compliance, analyzable signal time, interval to arrhythmia detection, and diagnostic yield, accuracy of arrhythmias detection, assess cardiac ectopy’s, the burden of arrhythmias and duration of arrhythmia, assess the accuracy of heart rate, P-wave detection and R-wave detection and evaluate the customer experience of the end-end solution Patch.

2 METHODS

The study was a single centre study conducted in all 51 volunteers for 24 hrs. Asymptomatic and symptomatic subjects with age 18 years and above. In general, exclusion criteria was the volunteers with comorbidity factors such as HTN, DM and previous CAD history, patients with Implantable Cardioverter Defibrillators (ICD) & pace makers, also who underwent chest-related surgeries within the last two months, pregnant and lactating women. Informed consent was taken from all participants and details of the study were explained to the participants in their local language. Institutional Ethics Committee (IEC) approved the study based on standard guidlines. The details of devices used was presented in Table 1.

Table 1: FDA and CE approved devices used for the study in comparision.

| Philips Digitrak XT | Vigo Smart Heart |
| --- | --- |
| Schiller Medilog FD5 plus | Vigo Smart Heart |
| Mortara H3+ | Vigo Smart Heart |
| RMS Cordis 3 | Vigo Smart Heart |

2.1 Working pipeline of VSH based monitoring system:

VSH (Vigo Holter) is simple patch plced on left side of the chest just below the clavical bone and does not need any additional glues for fixing. It is a single patient use, with two leads, wireless, disposable patch with biocompatibel adhesive, connected to cloud and internet of things (IoT) platform along with domain expert surveillance. It was designed as an easy wearable for the patient with user friendly instructions.the patient can engage in his regular activitiesand it records continuous ECG and HR. The VSH App was installed in the consumer‘s smart phone inpost monitoring period, data was uploaded to the secure central cloud location. In the cloud, the patient data was analysed by advanced Artificial Intelligence (AI engine) and manually validated by a qualified ECG technician. Final report was made available for physician/cardiologist to review and a copy was sent to the patient as well.

Fig. 1: Schematic representation of the working pipeline of VSH based monitoring system.

For the study Individual study case records were prepared for the participants. The collected data was computerised at central location and the repository was safely stored complying with the local data protection laws. The interventions with sufficient details to allow replication, including how and when they were actually administered were taken care.

- VSH and traditional holter monitor was attached to all the participants in the study.

- Recording period was 24 hours.

- Dairy of events to be recorded by the participants.

- Final recording ECG data analysed ECG report was sent to two different ECG reviewers and ECG was analysed blindfolded.

- Dairy of events were shared with the reviewers.

- Once the report is generated, the report and the data was further cross-verified by and independent Electrophysiologist (Blindfolded)

2.2 Un biased approach for analysis:

- ECG data and ECG report reviewer were randomly allocated.

- ECG report was generated by analysis softwares, ECG Holter reports were

randomly allocated to Electrophysiologist

2.3 Statistical analysis

- Trial limitations, addressing sources of potential bias, imprecision, and, if

relevant, multiplicity of analyses

- Cross-sectional study was conducted in limited time frame.

- Generalisability (external validity, applicability) of the trial findings

- Generalisations to be logically derived based on the results.

- Interpretation consistent with results, balancing benefits and harms, and

considering other relevant evidence

3 RESULTS

The total number of participants and detailed analysis based on demographics, diseases, symptoms etc. depicted in Table.2 . Very importantly Vigo smartheartidentified few major conditions and clinical events useful for  identification of  CVDs which were missed out by traditional holters. Results of data analyses were performed for both traditional and VSH including subgroup analyses and adjusted analyses, distinguishing pre-specified from exploratory. Below arefew important cases studies, where the results of ECG strips compared with traditional holter data presented in Table 3.

Table 2: Table showing number of participants and details analysis of results.

| Total number participants | 51 |
| --- | --- |
| Reports with equivalent result | 45 |
| Conditions only identified by traditional holter but not by Vigo Holter | 0 |
| Conditions only identified by Vigo reports but not by traditional Holter | 6 |
| Details of conditions that were not identified by traditional holter but identified by Vigo Holter | 1.Second Degree Mobitz Type I block in one patient 2. First Degree AV block in one patient 3. IVCD in two patients 4. SVT episodes in two patients |
| ECG wave quality in reports with differences | Same in both recordings |
| Total Diagnostic Yield | 39% |

Table 3: comparision of ECG data quality between Traditional holter and Vigo smartheart

|  | Traditional Holter | Vigo Smart Heart |
| --- | --- | --- |
| P Wave Quality | ✓ | ✓ |
| QRS Wave Quality | ✓ | ✓ |
| R-R Interval | ✓ | ✓ |
| P-R Interval | ✓ | ✓ |
| Total Noise | 1301 Minutes | 990 minutes |
| Total Analyzable time | 90.22% | 99.3% |

The output data of traditional holter was comparable to the data of Vigo SmartHeart. The P wave quality is same for both the holters. QRS Wave qualities, R-R Interval, P-R interval are also same which are very crucial for determining the myocardial arrhythmia. The total noise was measures as 1301 Minutes and 990 with traditional holter and VSH respectively. Out of total noise the analyzable time was exceptionally good with 99.3 % for VSH where as it is only 90.22% with traditional holter. The larger analyzable data enables better visibility to the findings made in ECG. For this clinical study, we have analyzed data for 51 patient subjects out of whom 45 subject’s VSH data was almost similar with Traditional holter ECG data quality.  The estimated net yielding capacity was 39%. There were six patients whose ECG findings were different.  The details of those cases are described below.

3.1 Case 1:

A 21-year-old male with no known past medical history presented to the emergency department with 3 days history of giddiness, exhaustion, fatigue and light-headedness but all the vital signs were normal. Patient advised for lab tests and also forprogressive prolongation of PR interval with non-conducted p wave with R-R Interval of 1.4seconds a long-term ECG along with VSH.   The Strip presented below where., Second degree AV block Mobitz I was identified by Vigo Smart Heart while traditional Holter report failed to mention it due to manual reading error.

Fig. 2: ECG strips comparison: Second degree AV block Mobitz ITraditional Holter Vs VSH.

3.2 Case 2:

A 55-year-old male, known case of HTN, DM and S/P PTCA admitted with the complaint of severe chest pain. The patient examined with a 12 lead ECG. The patient still had pain pointed in nature towards to the left chest wall. So the cardiologist advised for a Holter monitoring. Simultaneously placed on VSH also for a comparison analysis. As shown in Figure.3. the Strip reveals that Vigo Smart heart identified Supraventricular tachycardia episode whereas a traditional Holter analysis failed to identify it due to noise. SVT is an abnormally fast heartbeat that can be lethal to a patient. In addition, it has many arrhythmias under which usually arise from atria or the AV node above the ventricles. This requires immediate attention.

Fig. 3: Case 1 ECG  strips Comparision for SVT of traditional Holter and Vigo Smart Heart.

3.3 Case 3:

A 71-year-old patient, presented to a cardiology clinic for evaluation of intermittent palpitations over the previous 3-4 months. These episodes were irregular in nature and lasted for approximately 10-15 minutes.  No other significant comorbidities or no other medication intake revealed. No family history of sudden cardiac death in the family. A 30-day event monitor performed formonths prior to presentation, which was negative. Vital signs and physical examination were normal. Therefore, he was advised Vigo smart heart monitoring. Vigo Smart heart identified IVCD episode whereas a traditional Holter analysis failed to identify it due to base line noise.   IVCD indicates intraventricular conduction delay; with intraventricular conduction delay being most strongly associated with an increased risk of arrhythmic death.

IVCD detection by VSH

Base line noise by Traditional holter

Fig. 4: ECG strips comparison for IVCD of Traditional Holter  and VSH

3.4 Case 4:

With the aid of VSH in another patient aged 55 years old, a first-degree AV block was identified who admitted due to severe chest pain with a history of HTN, DM and S/P PTCA. The 12 lead ECG examination was unable to detect the abnormality or SVT in the patient.

Missed out by​

Traditional Holter​

Ist degree AV Block identified by VSH

Fig. 5: ECG strips comparison for Ist degree AV of Traditional Holter  and VSH

Noises present in the ECG signal that of high frequency such as electromyogram noise, additive white gaussian noise, power line interference  and baseline wandering which is of low frequency can be contaminated in the ECG signal leading to wrong interpretation. Highly predicatble analysable time provides the data that aids in approaching the faster clinical decisions and also appropriate management.  From VSH reports proved that no lead detachments and noise resultant from the wire entanglements leading to low noise and highly analyzable time.

4 DISCUSSION

Cardiovascular Disease related deaths are significantly increasing especially among younger generation due to life style, food habits, and individual genetics make up (Anand et al., 2016)(1). Amongst all the cardiac conditions, arrythmias remains the least addressed one. However, Arrhythmia situation of abnormal heart beat that can lead to irreparable damage or even death (Arevalo et al., 2016)(2).. Arrythmia detection is tradianally considered as combustion due to the presence to lots of wires. Thus the focus has to shift towards looking into more convinient option to detect arrythmias.  Vigo SmartHeart is one of such options which was measured against the traditional holters and has showed positive results.  In the present study we presented four case studies where Vigo holter detected important cardiac parameters including Second Degree Mobitz Type I block,  First Degree AV block ,  IVCD and  SVT episodes which are not detected by traditional holter.

Out of 51 subjects, 6 cases were identified for few major clinical conditions of  CVDs by VSH but failed by traditional Holter. Among the study subjects, second degree AV block Mobitz I was identified by VSH in 21-year-old male subject. When long-term ECG along with VSH was taken progressive prolongation of PR interval with non-conducted p wave with R-R Interval of 1.4 seconds was identified (Fig.2), as traditional Holter report failed to mention it due to manual reading error.  In few cases, traditional Holter analysis failed to identify SVT  due to its noise, where as Vigo Smart heart identified abnormal heartbeat (fig.3).  As shown Figure.4. Vigo Smart heart identified IVCD, involved and strongly associated with an increased risk of arrhythmic death. A traditional Holter analysis failed to identify it due to base line noise.

According to Cardioloists, The ECG is considered as the “gold standard” technique for diagnosis of cardiac arrhythmias, conduction disturbances, myocardial infarction with Q waves. Interpretation of ECG waves and singals are cruicla for diagnosis and tailoring the therapy. Nosies araising ECG signals such as high frequency such as electromyogram noise, additive white gaussian noise, power line interference  and baseline wandering which is of low frequency can be contaminated in the ECG signal leading to wrong interpretation. The outcome of the study showed that highly predicatble & analysable time of VSH provides the data helps clinicians to take the faster clinical decisions and also appropriate management and also that no lead detachments and noise resultant from the wire entanglements leading to low noise and highly analyzable time.

According to pulmonolgist,  the monitoring of vitals has become more easy just a click away and life becomes easy for the nursing staff to monitor a patient. In emergency situations they can take ontime and quick decisios before the doctor reaches the patient and save the patient to fall into life threatening situations by monitoring and maintating the vitals. Especially, during the post ICU monitoring period the trending visual graphs helps monitor the patient effectively because for every 5 minutes the VIGO application send updates to the patient, doctor and to the central  monitoring station.The patient will be with peace of mind since they know theirs vitals and health condition by time to time. During the COVID-19 pandemic it is better to minimal physical interaction with the patient. In this context the remote health monitoring platforms are very useful and life saving. The advantage of the platform is that at a time a doctor can monitor multiple patients through doctor interface dashboard in application so that effective management of patients becomes easy.

The patient aged 64 years from Hyderabad, India says the no doctor visits you when one is in home quarantine. VSH monitoring helps the patient and the doctor to know about the vital parameters of the patient to avoid serious complications, very useful for elderly people. Therefore, it is very useful during the pandemics like COVID-19.

6 CONCLUSION

The Vigo Smartheart demonstraed to have more clinical efficiency than the traditional holters. The testimonials from the users also established that the solution is easy to use and is completely wireless. The patient and family will be with peace of mind since they know their vitals just on their fingertips.   The study threw insights  that VSH can detect the conditions which were missed out by traditional holter monitoring. With its ease in using it can also be used  to detect and screen arrhythmias in healthy individuals, to differentiate cardiac and non cardiac origin syncope, palpitations, dizziness etc, Better patient compliance was noticed almost 100% patients completed  monitoring  without any inconvinene vs traditional holters. VSH is very much needed for CVDs and especilly during this COVID-19 pandemics for the patients in home quarantine.

FUNDING INFORMATION

Funding: Vigo Health Care

CONFLICT OF INTEREST

Authors have no conflict of interest.

REFERENCES

[1]	Babapoor-Farrokhran S, Rasekhi RT, Gill D, Babapoor S, Amanullah A. Arrhythmia in COVID-19. SN Compr Clin Med 2020;2:1430–5. https://doi.org/10.1007/s42399-020-00454-2.

[2]	Galli A, Ambrosini F, Lombardi F. Holter monitoring and loop recorders: From research to clinical practice. Arrhythmia Electrophysiol Rev 2016;5:136–43. https://doi.org/10.15420/AER.2016.17.2.

[3]	Anand SS, Hawkes C, De Souza RJ, Mente A, Dehghan M, Nugent R, et al. Food Consumption and its Impact on Cardiovascular Disease: Importance of Solutions Focused on the Globalized Food System A Report from the Workshop Convened by the World Heart Federation. J Am Coll Cardiol 2015;66:1590–614. https://doi.org/10.1016/j.jacc.2015.07.050.

[4]	Arevalo HJ, Vadakkumpadan F, Guallar E, Jebb A, Malamas P, Wu KC, et al. Arrhythmia risk stratification of patients after myocardial infarction using personalized heart models. Nat Commun 2016;7. https://doi.org/10.1038/ncomms11437.

[5]	Kumar P, Sharma VK. Detection and classification of ECG noises using decomposition on mixed codebook for quality analysis. Healthc Technol Lett 2020;7:18–24. https://doi.org/10.1049/htl.2019.0096.

[6]	E. Hoefman, P.J.E. Bindels HCPM van W. Efficacy of diagnostic tools for detecting cardiac arrhythmias: systematic literature search. n.d.

[7]	Ahmad F, Sneineh A, Patel RS, Reddy R, Llukmani A, Hashim A, et al. In The Line of Treatment: A Systematic Review of Paroxysmal Supraventricular Tachycardia 2021. https://doi.org/10.7759/cureus.15502.

[8]	Sisti N, Valente S, Mandoli GE, Santoro C, Sciaccaluga C, Franchi F, et al. COVID-19 in patients with heart failure: the new and the old epidemic n.d. https://doi.org/10.1136/postgradmedj-2020-138080.

[9]   Murat F, Yildirim O, Talo M, Baloglu UB, Demir Y, Acharya UR. Application of deep learning techniques for heartbeats detection using ECG signals-analysis and review. Comput Biol Med. 2020 May;120:103726. doi: 10.1016/j.compbiomed.2020.103726. Epub 2020 Apr 8. PMID: 32421643.

[10] Raj, S., Ray, K.C. A Personalized Arrhythmia Monitoring Platform. Sci Rep 8, 11395 (2018). https://doi.org/10.1038/s41598-018-29690-2.

[11] Hong S, Zhou Y, Shang J, Xiao C, Sun J. Opportunities and challenges of deep learning methods for electrocardiogram data: A systematic review. Comput Biol Med. 2020 Jul;122:103801. doi: 10.1016/j.compbiomed.2020.103801. Epub 2020 Jun 7. PMID: 32658725.

[12] Hoefman E, Bindels PJ, van Weert HC. Efficacy of diagnostic tools for detecting cardiac arrhythmias: systematic literature search. Neth Heart J. 2010;18(11):543-551. doi:10.1007/s12471-010-0831-0.