Validation of sodC gene-based PCR assay and antimicrobial resistance profiling of Neisseria meningitidis in asymptomatic carriers

Mekonnen Atimew1,2*, Melaku Yidenekachew1, Marchegn Yimer1, Ashenafi Alemu1, Dawit Hailu1, Tadelo Wondimagegn3, Fitsumbiran tajebe3, Gashaw Adane3, Tesfaye Gelanew1, Getachew Tesfaye Beyene1

Affiliations

1Armauer Hansen Research Institute (AHRI), Addis Ababa, Ethiopia

2 Department of Medical Laboratory Sciences, College of Medicine & Health Sciences, Wolaita Sodo University, Wolaita Sodo, South Ethiopia

3Department of Immunology and Molecular Biology, College of Medicine & Health Sciences, University of Gondar, Gondar Ethiopia

*Corrosponding authors:

Getachew Tesfaye Beyene - email. getachew.tesfaye@ahri.gov.et

Mekonnen Atimew - email: mekonnenatimew5@gmail.com

Abstract

Background: Neisseria meningitidis is one of the leading causes of bacterial meningitis and septicemia worldwide. The bacteriological culture was a widely used method for the detection of meningococcus, but it has low sensitivity and long waiting periods. Molecular detection targeting capsule transport gene was used, but over 16% of meningococcal carriage isolates lack ctrA and generate false-negative results due to sequence variations. The Cu-Zn superoxide dismutase gene (sodC) is specific to N. meningitidis, not found in other Neisseria species, making it better able to identify encapsulated meningococci and useful for detecting non-group-able meningococci without intact ctrA.

Objective: The objective of this study was the validation of sodC gene-based PCR assay and antimicrobial resistance profiling of N. meningitidis in asymptomatic carriers

Methods: The sodC gene N. meningitidis detection method was developed using a pair of primers and optimized. A total of 137 archived samples that were collected from the asymptomatic carrier suspected of having meningococcal infection were used for validation of the assay. STATA version 14.0 was used for analysis of clinical and demographic data after the data was entered into Epi Info version 7. Graphs and frequency tables from descriptive statistics were used to summarize the outcome. Two-by-two tables were used to compare the sensitivity specifically between the sodC-based PCR assay and culture and ctrA-based PCR.  A Disk diffusion test was used to determine the antimicrobial sensitivity of the isolates against antimicrobial drugs. To determine the association between independent and outcome variables, bivariate and multivariate logistic regression models were used and P-values less than 0.05 were considered statistically significant.

Result: The PCR assay targeting the sodC gene detected N. meningitidis DNA in 105 (76.6%) out of 137 clinical samples, while ctrA-based PCR detected the pathogen in 64 (46.7%) of the samples, and 49 samples (35.8%) of N. meningitidis were identified by culture. Then, the concordance of our in-house PCR assay targeting the sodC gene with ctrA PCR was performed using 137 clinical samples (Nasopharyngeal swabs). Among the 49 DNA samples from culture-positive N. meningitidis isolates used for validation, the sodC gene-based PCR accurately identified all 49 culture-confirmed isolates. In contrast, the ctrA gene-based PCR detected only 33 of these isolates. Out of the 49 N. meningitidis isolates by culture 43 (87.8%), 42 (83.7%), 32 (65.3%), 22 (44.9%), and 18 (36.7%), and 7 (15.2%) were resistant to amoxicillin, ampicillin, trimethoprim-sulfamethoxazole, ceftazidime, ceftriaxone, and meropenem, respectively. Furthermore, the majority of N. meningitidis isolates 36 (73.5%) were sensitive to cefepime, 31 (63.3%) of them were sensitive to ceftriaxone and meropenem, and 26 (53.1%) of them were sensitive to ceftazidime.

Conclusion: The sodC gene-based PCR assay demonstrated high sensitivity in detecting N. meningitidis in carriage specimens compared to ctrA gene-based PCR. The high prevalence of antibiotic resistance observed is alarming and highlights the urgent need to continue monitoring antibiotic resistance to inform treatment strategies effectively.

Keywords:  PCR, Molecular diagnostics, Neisseria meningitidis, sodC, ctrA, Antimicrobial resistance, Ethiopia

Introduction

Neisseria meningitidis, often referred to as meningococcus, is a gram-negative bacterium that can cause a spectrum of diseases ranging from mild sepsis with rapid recovery to fulminant meningococcemia (1). They are carried by healthy individuals in the nasopharynx around 5-10% (2). Every year, more than 1.2 million cases of bacterial meningitis are estimated to occur globally, causing about 14.25% of deaths (3). In Ethiopia, bacterial meningitis is an important cause of premature death and disability, being the 9th most common cause of years of life lost and disability-adjusted life years. Early treatment is essential in the clinical management of meningitis. A delay in therapy negatively affects the prognosis for patients with meningitis (4). The incidence of meningitis varies depending on age, geographic location, species, genotype, strain, and serotype of the causative agents (5, 6).

Accurate diagnosis and early treatment of meningococcal infections are essential due to their worldwide distribution, increased case fatality and morbidity rate, epidemic potential, and serious complications that can occur (7). For confirming the etiology, CSF and/or blood culture have been used as the gold standard for the diagnosis of meningococcal infection (8). Identification and detection of bacterial pathogens in cases of suspected meningitis are important in guiding appropriate treatment and prophylaxis. However, diagnosis of bacterial meningitis is often difficult (9). Traditional laboratory diagnosis of meningococcal disease (MD) has relied heavily on bacteriologic culture methods, but the bacterial growth rates, particularly in patients who have received pre-admission antibiotic treatment, are very low and have low sensitivity due to the frequent initiation of antimicrobial therapy before clinical sample collection (10). Many studies have shown that the high rates of morbidity and mortality linked to MD, particularly in children, are partly caused by delayed detection and diagnosis (11-14).

Molecular assays can be used for the diagnosis of invasive meningococcal infections when previous antibiotic therapy may inhibit bacterial growth (8). In previous reports, either ctrA or sodC genes were independently used for detecting meningococcal DNA, which has different sensitivities, both of which are superior to culture (15-18). The genome of N. meningitidis displays extensive genetic variability within a single course of infection by genome rearrangement (19).  The high genetic diversity of N. meningitidis makes reliable laboratory detection of this important pathogen difficult (20-22). A ctrA was thought to be found in all invasive strains of N. meningitidis due to the importance of the capsule in preventing complement-mediated killing (23). However, there has been a question raised about the sensitivity of ctrA in the identification of invasive N. meningitidis isolates. Several cases of invasive and sometimes fatal disease caused by capsule-null (cnl) strains lacking ctrA have been described (24-27).

Given that ctrA is not present in 16% or more of carriage isolates, it is not a suitable target gene for PCR on carriage study specimens. The reliability of ctrA-based PCR assays in detecting non-groupable invasive N. meningitidis isolates or specimens containing them is in question chiefly due to the genetics of non-groupable N. meningitidis carriage isolates and their genome plasticity. Because of natural competence, frequent horizontal gene transfer, and high rates of recombination, N. meningitidis has a significant degree of genomic diversity, which makes it difficult to utilize a single gene for identification (16, 28, 29). Also, ctrA-based molecular assay has been reported to generate false-negative results due to sequence variations in ctrA (15, 28, 30). Non-capsular genes like sodC, metA, and tauE provide complementary targets to ctrA to improve the detection of N. meningitidis via multiplex PCR (31).

Early antibiotic treatment of meningococcal disease is crucial for keeping the case fatality rate and risk of sequelae as low as possible (32). Resistance to various antimicrobial drugs used either for therapy of invasive infections or for prophylaxis of case contacts has long been recognized. There is limited data on the antimicrobial susceptibility test (AST) pattern of N. meningitidis in Ethiopia in general (27, 33).

The development of reliable molecular approaches for the detection and characterization of N. meningitidis is clinically relevant. Molecular methods offer several advantages over culture-based methods, including increased sensitivity, specificity, speed, and efficiency. Furthermore, generating information about the AST pattern of N. meningitidis is crucial to identify resistance in the meningococcus to prevent treatment failure because of the administration of inappropriate antibiotics (34).

Thus, the aims of this study were validation of sodC gene-based PCR assay and antimicrobial susceptibility pattern of N. meningitidis isolates.

MATERIALS AND METHODS

Ethical consideration

This study was done on clinical (pharyngeal swabs) and culture samples that were collected from a MenAfricar study in 2016 that employed a cross-sectional study and stored in the AHRI. Ethical waiver to use those archived isolates and clinical samples from All African Lepersoy Rehabilitation Center (ALERT)/ Armauer Hansen Research Institute (AHRI) Ethics Committee with approval number: PO-63-22.

Study Isolates and Clinical Samples

In 2016, respiratory clinical specimens were collected from 3000 asymptomatic Neisseria spp carriers. The leftover specimens (pharyngeal swabs) and culture-positive, 49 isolates, were stored at -80°C in 1 ml of STGG medium (containing Skim milk, Tryptisoya, Glucose, and Glycerol) at the AHRI microbiology laboratory. To assess its performance in clinical specimens, 137 archived specimens preserved in 1 ml of STGG were randomly selected and the sodC gene-based PCR assay was validated.  Demographically of the total clinical samples, 71 (52%) were collected from males and 66 (48%) were from females. The age categories of 15–19 had the highest percentage (26%) while the age groups of 1-4 had the lowest (6%). The age group ranges from 1 to 29 years, with a mean age of 16.38 years (Table 1).

Characterization and confirmation of N. meningitidis isolates

Cryopreserved bacterial isolates were thawed at room temperature. A loopful of these samples was then sub-cultured onto fresh chocolate (CHO) and modified Thayer Martin (MTM) agar plates, which were supplemented with vancomycin, colistin, nystatin, trimethoprim (VCNT), and IsoVitaleX enrichment. Colonies were further confirmed using gram stain and oxidase tests. After confirming the presence of gram-negative diplococci and oxidase-positive isolates, a pure colony was sub-cultured again on a blood agar plate (BAP) with 5-10% CO2 for 24 to 48 hours to ensure the purity of the culture for further biochemical testing.

A carbohydrate utilization test (glucose, maltose, lactose, and sucrose) was performed using cystine trypticase agar (CTA) to differentiate N. meningitidis from Moraxella species and other nonpathogenic Neisseria species. Isolates that were gram-negative diplococci, oxidase-positive, glucose fermenters, maltose fermenters, and non-fermenters of lactose and sucrose were confirmed as N. meningitidis, a total of 49 isolates.

Antimicrobial susceptibility testing (AST)

AST was performed on 49 isolates of N. meningitidis, all of which exhibited the standard characteristics of N. meningitidis colonies using Mueller-Hinton agar (MHA) with 5% sheep blood (34). Briefly, 3-5 colonies from a blood agar plate were suspended, and the turbidity was adjusted to match a 0.5 McFarland standard. The surface of the MHA with 5% sheep blood agar was then thoroughly coated with the bacterial suspension using a sterile swab. After allowing the plate to dry for 3-5 minutes, antibiotic discs were evenly placed on the inoculated plate with sterile forceps. The plate was incubated in 5% CO2 at 37ºC for 24 hours. The antibiotics selected for AST profiling were based on the Ministry of Health Ethiopia guidelines, fourth edition 2021(https://www.slideshare.net/TesfayeWorkie/stg-2021pdf#1). The tested antibiotics included Ceftriaxone (30μg), Meropenem (30µg), Cefepime (30µg), Trimethoprim-sulfamethoxazole (1.25/23.75µg), Ampicillin (10µg), Amoxicillin (10µg), and Ceftazidime (30µg). The diameters of the inhibition zones around the discs were measured to the nearest millimeter using a graduated caliper, and these measurements were compared to standard charts to determine the susceptibility, intermediate resistance, or resistance of the bacteria to the tested antibiotics, according to clinical and laboratory standard institute 2022.

DNA Preparation and Quantification

DNA was isolated from 137 clinical samples (pharyngeal swabs) preserved in STGG broth, as well as from the 49 isolates. Briefly, samples stored in 1 ml of STGG broth at -80ºC were thawed at room temperature. Then, 200 μl of the sample was mixed with 200 μl of 1x TE buffer; in the case of the cultured sample, 3-5 colonies were picked up from the agar plate, then vortexed vigorously, and incubated in a water bath at 95ºC for 15 minutes. The tubes were then transferred to -20ºC for 10 minutes to release bacterial DNA through heat shock, followed by a 2-minute incubation at room temperature. The tubes were centrifuged at 14,000 rpm for 5 minutes, and the genomic DNA remained in the upper aqueous phase (supernatant). The supernatant containing the DNA was transferred to a separate 2 ml sterile tube, and the DNA concentration was measured using a NanoDrop 2000/2000c spectrophotometer (Thermo Fisher Scientific, USA). The DNA was then stored at -80ºC until used as a template for the PCR assay.

In-house Development of sodC-based PCR assay

We selected the sodC gene as the target for developing a PCR assay aimed at detecting N. meningitidis in clinical samples. The sodC gene was chosen for several reasons: (i) it offers high specificity for detecting N. meningitidis as it is absent in other Neisseria species; (ii) the sodC gene in N. meningitidis encodes virulence factors and a periplasmic enzyme, making it less prone to antigenic variation due to selective pressure and (iii) there are no known strains of meningococci that lack the sodC gene. These factors collectively suggest that a PCR assay based on the sodC gene can detect all N. meningitidis strains from various geographical regions without cross-reacting with other Neisseria species (15).

Primer design

Primers (both forward and reverse) were designed using SnapGene Viewer (Table 2). The sodC sequences of N. meningitidis were sourced from Gene Bank (accession number >NZ_CP021520.1:999157-999717). The specificity of these primers to N. meningitidis was verified using the NCBI nucleotide BLAST tool.

sodC PCR amplification conditions

The sodC PCR reaction was performed in a 25 μl mixture, which included 0.625 μl of each primer (10 μM), 2.5 μl of dNTPs (2.5 μM), 0.5 μl of DNA polymerase, 2.5 μl of polymerase buffer with magnesium acetate (5 μl of 10x), 15.25 μl of nuclease-free water, and 3 μl of genomic DNA as the template. Each run incorporated both positive and negative controls. The optimized amplification protocol started with an initial denaturation at 95°C for 3 minutes, followed by 35 cycles of denaturation at 94°C for 10 seconds, annealing at 56°C for 30 seconds, extension at 72°C for 30 seconds, and a final extension at 72°C for 10 minutes. The amplified product was then analyzed on a 2% agarose gel using 1x Tris-acetate EDTA (TAE) buffer and visualized with Bio-Rad UV scanning (a chemi PRO UV scanner, USA). The optimal conditions for the sodC PCR reaction and the most effective primer pairs were identified. This assay was then validated using the culture technique as the gold standard and then compared with the ctrA PCR assay.

ctrA gene-based PCR assay

The ctrA gene-based PCR assay, which includes the paired oligonucleotide primer (Table 3), was initially developed by Lansac N, et.a1 in 2000 (35), and later used at AHRI (36). Briefly, the PCR reactions comprised: 0.625 μl of both forward and reverse primers (10 μM each), 2.5 μl of dNTPs (2.5 μM), 0.5 μl of DNA polymerase, 2.5 μl of polymerase buffer with magnesium acetate (5 μl of 10x), 15.25 μl of nuclease-free water, and 3 μl of genomic DNA as the template, making a final volume of 25 μl. Each run included positive and negative controls. The amplification protocol was optimized with an initial denaturation at 95°C for 3 minutes, followed by 30 seconds at 94°C for denaturation, 45 seconds at 52°C for annealing, and 1 minute at 72°C for extension, with a final extension at 72°C for 10 minutes over 35 cycles. The amplified product was analyzed on a 2% agarose gel using 1x TAE buffer and visualized with Bio-Rad UV scanning (a chemi PRO UV scanner, USA). Representative gel images of the PCR products from both the sodC and ctrA genes, obtained from control strains ATCC (serogroup A: Z2491; W: A22; X: 860060; Y: 71/94) and suspected samples, are shown in Figure 2.

Data Management and Analysis

The data obtained was entered into Epi Info version 7 and exported into STATA version 14 for data cleaning and analysis. Results were presented in frequency tables and charts, descriptively. Bivariate and multivariate logistic regression models were used to analyze data and P-values less than 0.05 were considered statistically significant.

Results

Optimal conditions for the sodC gene-based PCR assay

To determine the optimal conditions for sodC gene-based PCR assay, a set of primer pairs targeting sodC gene was optimized by using the temperature gradient of the annealing temperature (Figure 1) and concentration of oligonucleotide forward and reverse primer pairs (Table 2). The sodC gene target was amplified using a selected optimized concentration of pair of primers at optimized PCR conditions: initial denaturation at 95°C for 3 min, followed by 35 cycles, denaturation at 94°C for 10 sec, annealing at 56°C for 30 sec, extension at 72°C for 30 sec, and a final extension at 72°C for 10 min.

Performance comparison between sodC gene-based detection of N. meningitidis and ctrA-based detection

To validate our in-house developed sodC gene-based assay for detecting N. meningitidis, we utilized culture-positive N. meningitidis isolates and compared the results with those from the ctrA-based PCR detection method. Among the 49 DNA samples from culture-positive N. meningitidis isolates used for validation, the sodC gene-based PCR accurately identified all 49 culture-confirmed isolates. In contrast, the ctrA gene-based PCR detected only 33 of these isolates. This demonstrates a 100% concordance between the culture method and the sodC gene-based PCR assay (Table 4). After validating our in-house PCR assay targeting the sodC gene with culture-positive isolates, we assessed its concordance with the PCR assay targeting the ctrA gene using 137 clinical pharyngeal swab samples. The sodC gene-based method detected N. meningitidis DNA in 76.64% of the samples, whereas the ctrA gene-based method identified it in only 46.72% (Table 5).

The bivariate and multivariate analyses revealed that variables such as sex, age, and residence of the study participants did not have a significant impact on the detection of N. meningitidis using the sodC gene-based PCR method (Table 6).

Antimicrobial susceptibility pattern of N. meningitidis isolates

The antimicrobial susceptibility patterns of the 49 N. meningitidis isolates are summarized in Table 7. Among these isolates, resistance was found in 43 (87.8%) to amoxicillin, 42 (83.7%) to ampicillin, 32 (65.3%) to trimethoprim-sulfamethoxazole, 22 (44.9%) to ceftazidime, and 18 (36.7%) to both ceftriaxone and meropenem. Additionally, 7 isolates (15.2%) showed resistance to cefepime. On the other hand, a notable number of isolates were sensitive to cefepime (36 isolates, 73.5%), ceftriaxone and meropenem (31 isolates, 63.3%), and ceftazidime (26 isolates, 53.1%).

Discussion

The results of this study offer important insights into the diagnostic techniques for the detection of N. meningitidis, particularly focusing on comparing PCR-targeted sodC and ctrA genes. The findings clearly show that, compared to the ctrA gene, the sodC-based PCR is a more sensitive and accurate method to detect N. meningitidis.

The higher detection rate of N. meningitidis using PCR targeting the sodC gene highlights the potential of this method as an effective tool for the diagnosis of meningococcal disease (37). In contrast, the lower detection rates of N. meningitidis using the ctrA gene methods suggest that this method may be less sensitive and reliable for detecting bacteria in clinical samples. This has important implications for clinical diagnosis, as timely and accurate detection of N. meningitidis is essential for initiating appropriate treatment and implementing public health measures to prevent the spread of the disease. A similar study that tested pharyngeal swabs for the presence of N. meningitidis by PCR with sodC and ctrA as target genes identified 75.8% (491/647) of clinical samples were tested positive for the sodC gene (38).

Furthermore, sodC-based real-time PCR identified a higher detection rate of N. meningitidis isolates, 518 of 520 (99.6%) isolates of N. meningitidis were positive for sodC, whereas ctrA detection occurs only in 368 of 520 (70.8%) isolates. Similar to our report, sodC gene-based PCR shows higher sensitivity than the ctrA gene (15).

Contrary to our finding, the same molecular assay for the detection of meningococci in normally sterile sites sodC and ctrA genes was used, and sodC-based RT-PCR was found less sensitive by 7.5% than ctrA (16).

Resistance to trimethoprim-sulphamethoxazole was high among meningococcal isolates. The widespread resistance to this antibiotic is possibly due to the early introduction of sulphonamides (39). In contrast, an investigation in children in Greece in 2004 revealed that ceftriaxone sensitivity was present in all isolates (40). There was no significant effect of the study participant's sex and age on the AST pattern of N. meningitidis (Table 8 and 9 respectively). However, we noted that the percentile of antibiotic resistance of N. meningitidis to different drugs differed slightly based on the age group; for example, those aged 24-29 showed the lowest percentile of resistance to ampicillin and amoxicillin, while N. meningitidis isolates from those aged between 10 and 14 showed the highest percentile of resistance to those of the same antimicrobial drugs.

N. meningitidis is developing a resistance to antibiotics that are recommended for the management of meningococcal meningitis for epidemic response. These findings have important implications for clinical practice and public health. Healthcare providers should be aware of the drug susceptibility profile of N. meningitidis in their region and consider these factors when selecting antimicrobial therapy for patients with suspected or confirmed meningococcal infections. Additionally, public health efforts should focus on surveillance of antimicrobial resistance in N. meningitidis and the development of new treatment strategies to combat resistant strains. Furthermore, the ability of PCR targeting the sodC gene to detect a larger proportion of N. meningitidis isolates has implications for epidemiological studies and surveillance of meningococcal disease. Accurate and comprehensive surveillance data are essential for understanding the epidemiology and dynamics of N. meningitidis transmission, as well as for informing public health interventions such as vaccination strategies. The higher detection rate of N. meningitidis using PCR targeting the sodC gene may also be valuable for identifying asymptomatic carriers of the bacteria, which is critical for implementing targeted control measures to prevent outbreaks.

Conclusion

The sodC gene-based PCR assay is found to be superior in sensitivity to detecting N. meningitidis in carriage specimens compared to ctrA gene-based PCR. The observed high prevalence of antibiotic resistance warrants the need to continue to monitor antibiotic resistance that might influence treatment and for future implementation of chemoprophylaxis for carriers and those household members who have contacts with confirmed meningitis cases.

Limitations of the study

The study was conducted in a limited number of samples archived in the laboratory.

Abbreviations

AHRI, Armauer Hansen Research Institute; AST, Antimicrobial susceptibility testing; ATCC, American Type Culture Collection; CSF, Cerebrospinal fluid; ctrA, Capsule transport to cell surface gene; DNA, Deoxyribonucleic acid; MD, Meningococcal disease; MHA, Muller Hinton Agar; PCR, Polymerase chain reaction; sodC, Cu-Zn superoxide dismutase gene

Data Sharing Statement

The data supporting the findings of this study are available upon reasonable request from the corresponding author.

Acknowledgments

I am grateful to the University of Gondar, the College of Medicine and Health Science, the School of Biomedical and Laboratory Sciences, the College of Medicine and Health Sciences, the Department of Immunology and Molecular Biology, and the Armauer Hansen Research Institute for their guidance and willingness to provide me with the opportunity to conduct research.

Funding

This work is supported by Grand Challenge Ethiopia and the African Academy of Sciences through the ARISE-PP project.

Disclosure

The authors declare no conflicts of interest.

References

1.	Manchanda V, Gupta S, Bhalla P. Meningococcal disease: history, epidemiology, pathogenesis, clinical manifestations, diagnosis, antimicrobial susceptibility and prevention. Indian journal of medical microbiology. 2006;24(1):7-19.

2.	Bårnes GK, Kristiansen PA, Beyene D, Workalemahu B, Fissiha P, Merdekios B, et al. Prevalence and epidemiology of meningococcal carriage in Southern Ethiopia prior to implementation of MenAfriVac, a conjugate vaccine. BMC infectious diseases. 2016;16:1-14.

3.	Bianchi L, Napoli Z, Donati S, Lencioni P, Santoni F, Lari R. Emergency management in bacterial meningitis and sepsis: application of real time-polymerase chain reaction and FilmArray technology performed directly on cerebrospinal fluid and blood samples. Microbiologia Medica. 2015;30(2).

4.	Bårnes GK, Gudina EK, Berhane M, Abdissa A, Tesfaw G, Abebe G, et al. New molecular tools for meningitis diagnostics in Ethiopia–a necessary step towards improving antimicrobial prescription. BMC infectious diseases. 2018;18(1):1-14.

5.	Başpınar EÖ, Dayan S, Bekçibaşı M, Tekin R, Ayaz C, Deveci Ö, et al. Comparison of culture and PCR methods in the diagnosis of bacterial meningitis. brazilian journal of microbiology. 2017;48:232-6.

6.	Rosenberg E, DeLong EF, Lory S, Stackebrandt E, Thompson F, Lory S. The prokaryotes: prokaryotic physiology and biochemistry: Springer; 2013.

7.	Nemescu RE, Ursu RG, Dorobăț CM, Iancu LS. The efficiency of sodC gene/N. meningitidis detection in comparison with the classical methods for the diagnosis of meningococcal infection/Evaluarea eficienţei Real Time PCR TaqMan utilizând gena sodC/N. meningitidis în comparaţie cu metodele clasice utilizate în diagnosticul infecţiei meningococice. Revista Romana de Medicina de Laborator. 2015;23(1):21-30.

8.	Qurbanalizadegan M, Ranjbar R, Ataee R, Hajia M, Goodarzi Z, Farshad S, et al. Specific PCR assay for rapid and direct detection of Neisseria meningitidis in cerebrospinal fluid specimens. Iranian journal of public health. 2010;39(4):45.

9.	Seward R, Towner K. Evaluation of a PCR-immunoassay technique for detection of Neisseria meningitidis in cerebrospinal fluid and peripheral blood. Journal of medical microbiology. 2000;49(5):451-6.

10.	Wylie PA, Stevens D, Drake W, Stuart J, Cartwright K. Epidemiology and clinical management of meningococcal disease in west Gloucestershire: retrospective, population based study. Bmj. 1997;315(7111):774-9.

11.	Kuppermann N, Malley R, Inkelis SH, Fleisher GR. Clinical and hematologic features do not reliably identify children with unsuspected meningococcal disease. Pediatrics. 1999;103(2):e20-e.

12.	Riordan FAI, Thomson AP, Sills JA, Hart CA. Lesson of the Week: Who spots the spots? Diagnosis and treatment of early meningococcal disease in children. Bmj. 1996;313(7067):1255-6.

13.	Sullivan TD, LaScolea Jr LJ. Neisseria meningitidis bacteremia in children: quantitation of bacteremia and spontaneous clinical recovery without antibiotic therapy. Pediatrics. 1987;80(1):63-7.

14.	BAUDIN C, CARUBA S, CLAESSENS Y. Diagnostic de pneumonie aiguë communautaire du sujet âgé aux urgences.

15.	Dolan Thomas J, Hatcher CP, Satterfield DA, Theodore MJ, Bach MC, Linscott KB, et al. sodC-based real-time PCR for detection of Neisseria meningitidis. PloS one. 2011;6(5):e19361.

16.	Higa FT, Fukasawa LO, Gonçalves MG, Salgado MM, Lemos APSd, Harrison LH, et al. Use of sodC versus ctrA for real-time polymerase chain reaction-based detection of Neisseria meningitidis in sterile body fluids. Memórias do Instituto Oswaldo Cruz. 2013;108:246-7.

17.	Abdeldaim GM, Strålin K, Korsgaard J, Blomberg J, Welinder-Olsson C, Herrmann B. Multiplex quantitative PCR for detection of lower respiratory tract infection and meningitis caused by Streptococcus pneumoniae, Haemophilus influenzae and Neisseria meningitidis. BMC microbiology. 2010;10(1):1-9.

18.	Azzari C, Canessa C, Lippi F, Moriondo M, Indolfi G, Nieddu F, et al. Distribution of invasive meningococcal B disease in Italian pediatric population: implications for vaccination timing. Vaccine. 2014;32(10):1187-91.

19.	Hwang S, Gou Z, Kuznetsov IB. DP-Bind: a web server for sequence-based prediction of DNA-binding residues in DNA-binding proteins. Bioinformatics. 2007;23(5):634-6.

20.	Patel R. MALDI-TOF MS for the diagnosis of infectious diseases. Clinical chemistry. 2015;61(1):100-11.

21.	Houlihan CF, Bharucha T, Breuer J. Advances in molecular diagnostic testing for central nervous system infections. Current opinion in infectious diseases. 2019;32(3):244-50.

22.	Unalan-Altintop T, Karagoz A, Hazirolan G. A diagnostic challenge in clinical laboratory: misidentification of Neisseria subflava as Neisseria meningitidis by MALDI-TOF MS. Acta microbiologica et immunologica Hungarica. 2020;67(4):258-60.

23.	Nolfi-Donegan D, Konar M, Vianzon V, MacNeil J, Cooper J, Lurie P, et al. Fatal nongroupable Neisseria meningitidis disease in vaccinated patient receiving eculizumab. Emerging Infectious Diseases. 2018;24(8):1561.

24.	Ganesh K, Allam M, Wolter N, Bratcher HB, Harrison OB, Lucidarme J, et al. Molecular characterization of invasive capsule null Neisseria meningitidis in South Africa. BMC microbiology. 2017;17(1):1-10.

25.	Johswich KO, Zhou J, Law DK, St. Michael F, McCaw SE, Jamieson FB, et al. Invasive potential of nonencapsulated disease isolates of Neisseria meningitidis. Infection and immunity. 2012;80(7):2346-53.

26.	Findlow H, Vogel U, Mueller JE, Curry A, Njanpop-Lafourcade B-M, Claus H, et al. Three cases of invasive meningococcal disease caused by a capsule null locus strain circulating among healthy carriers in Burkina Faso. The Journal of infectious diseases. 2007:1071-7.

27.	Hoang LM, Thomas E, Tyler S, Pollard AJ, Stephens G, Gustafson L, et al. Rapid and fatal meningococcal disease due to a strain of Neisseria meningitidis containing the capsule null locus. Clinical Infectious Diseases. 2005;40(5):e38-e42.

28.	Jaton K, Ninet B, Bille J, Greub G. False-negative PCR result due to gene polymorphism: the example of Neisseria meningitidis. Journal of clinical microbiology. 2010;48(12):4590-1.

29.	Brik A, Terrade A, Hong E, Deghmane A, Taha M-K, Bouafsoun A, et al. Phenotypic and genotypic characterization of meningococcal isolates in Tunis, Tunisia: High diversity and impact on vaccination strategies. International Journal of Infectious Diseases. 2020;91:73-8.

30.	Cavrini F, Liguori G, Andreoli A, Sambri V. Multiple nucleotide substitutions in the Neisseria meningitidis serogroup C ctrA gene cause false-negative detection by real-time PCR. Journal of clinical microbiology. 2010;48(8):3016-8.

31.	Sirluck-Schroeder I, Al-Rawahi GN, Gadkar V, Hoang L, Tsang R, Tilley P. Limitation of ctrA as a Target for Neisseria meningitidis Identification and Potential Alternative Targets. Journal of Clinical Microbiology. 2022;60(6):e00152-22.

32.	Hedberg ST, Fredlund H, Nicolas P, Caugant DA, Olcén P, Unemo M. Antibiotic susceptibility and characteristics of Neisseria meningitidis isolates from the African meningitis belt, 2000 to 2006: phenotypic and genotypic perspectives. Antimicrobial agents and chemotherapy. 2009;53(4):1561-6.

33.	Mulatu F, Yeshitela B, Tilahun H, Yidnekachew M, Yimer M. Antibiotic susceptibility patterns of neisseria meningitides isolates from asymptomatic carriers in Gurage Zone, Southern Ethiopia. Am J Health Res. 2019;7(1):12-8.

34.	Taha M-K. Molecular detection and characterization of Neisseria meningitidis. Expert Review of Molecular Diagnostics. 2002;2(2):143-50.

35.	Lansac N, Picard F, Menard C, Boissinot M, Ouellette M, Roy P, et al. Novel genus-specific PCR-based assays for rapid identification of Neisseria species and Neisseria meningitidis. European journal of clinical microbiology and infectious diseases. 2000;19:443-51.

36.	Diallo K, Coulibaly MD, Rebbetts LS, Harrison OB, Lucidarme J, Gamougam K, et al. Development of a PCR algorithm to detect and characterize Neisseria meningitidis carriage isolates in the African meningitis belt. PLoS One. 2018;13(12):e0206453.

37.	Harrison LH. Prospects for vaccine prevention of meningococcal infection. Clinical microbiology reviews. 2006;19(1):142-64.

38.	Morselli S, Gaspari V, Cantiani A, Salvo M, Foschi C, Lazzarotto T, et al. Meningococcal carriage in ‘men having sex with men’with pharyngeal gonorrhoea. Frontiers in Cellular and Infection Microbiology. 2022;11:1400.

39.	Rohani M, Ahmad Afkhar F, AL AM, Muhd Amir K, Sahura H, Fairuz A, et al. Serogroups and antibiotic susceptibility patterns of Neisseria meningitidis isolated from army recruits in a training camp. Malaysian Journal of Pathology. 2008;29(2):91-4.

40.	Pavlopoulou I, Daikos G, Alexandrou H, Petridou E, Pangalis A, Theodoridou M, et al. Carriage of Neisseria meningitidis by Greek children: risk factors and strain characteristics. Clinical microbiology and infection. 2004;10(2):137-42.

Table 1: Socio-demographic characteristics for the asymptomatic nasopharyngeal carriage rate of N. meningitidis

| Characteristics of the study population (n=137) |  | Number | Percentage |
| --- | --- | --- | --- |
| Sex | Male | 71 | 52 |
|  | Female | 66 | 48 |
| Residence | Urban | 58 | 42.34 |
|  | Rural | 79 | 57.66 |
| Age | 1-4 | 8 | 6 |
|  | 5-9 | 15 | 11 |
|  | 10-14 | 33 | 24 |
|  | 15-19 | 36 | 26 |
|  | 20-24 | 31 | 23 |
|  | 25-29 | 14 | 10 |

Table 2: List of Oligonucleotide primers designed for the N. meningitidis sodC gene target

| S. no | Oligonucleotide | 5’-3’ nucleotide sequences |
| --- | --- | --- |
| 1. | sodC Fw1-PCR | ATGAATATGAAAACCTTATTAGCACTAGCGGTTAGTGCAG |
| 2. | sodC Fw14-48 | CCTTATTAGCACTAGCGGTTAGTGCAGTATGTTC |
| 3. | sodC Fw14-PCR | CCTTATTAGCACTAGCGGTTAG |
| 4. | sodC Fw64 | GCACACGAGCATAATACGATACCTAAAGGTGCTTC |
| 5. | sodC Fw118 | CAACTTGATCCAGCAAACGGTAACAAAGATGTGGG |
| 6. | sodC Fw361 | GCACACTTAGGTGATTTACCTGCATTAACTG |
| 7. | sodC Rv478-PCR | GGATCATAATAGAGTGACCGCGAAC |
| 8. | sodC Rv520-PCR | CAAGTGGAGCTGGATGATCGGAGTG |
| 9. | sodC Rv561-PCR | TTATTTAATCACGCCACATGCCATACGTGG |

Table 3: Oligonucleotide primers used in this study for the ctrA gene target

| Gene and group | Designation | 5’-3’ nucleotide sequences | Bp |
| --- | --- | --- | --- |
| CtrA | ctrA-F | GCTGCGGTAGGTGGTTCAA | 110 |
|  | ctrA-R | TTGTCGCGGATTTGCAACTA |  |

Table 4: Comparative concordance of the two molecular detection techniques on 49 N. meningitidis isolates confirmed by culture

|  | Positive | Negative | Total |
| --- | --- | --- | --- |
| sodC | 49 | 0 | 49 |
| ctrA | 33 | 16 | 49 |

Table 5: Comparison of the sodC gene-based PCR assay and the ctrA gene-based PCR assay for detecting N. meningitidis in 137 DNA samples derived from 137 clinical specimens

|  | N (DNA samples) | Positive | Negative | Total |
| --- | --- | --- | --- | --- |
| sodC | 137 | 105 | 32 | 137 |
| ctrA | 137 | 64 | 73 | 137 |

Table 6: Bivariable and multivariable logistic regression analysis for independent variables associated with N. meningitidis DNA detected using sodC gene-based PCR

| Independent Variables | SodC gene-based PCR |  | OR (95% CI) | p-value | AOR (95% CI) | p-value |
| --- | --- | --- | --- | --- | --- | --- |
|  | Positive | Negative |  |  |  |  |
| Sex Male Female | 37 15 | 195 209 | 1.1 (0.5, 2.4) 1.0 | 0.81 | 1.04 (0.47, 2.32) 1.0 | 0.93 |
| Age ≥21 11-20 <10 | 29 57 19 | 9 15 8 | 1.4 (0.45, 4.13) 1.6 (0.59, 4.36) 1.0 | 0.59 0.36 | 1.32 (0.43, 4.05) 1.52 (0.55, 4.19) 1.0 | 0.63 0.42 |
| Resident Urban Rural | 43 62 | 10 22 | 1.5 (0.65, 3.5) 1.0 | 0.33 | 1.47 (0.63, 3.44) 1.0 | 0.11 |

Table 7: Drug susceptibility profile of 49 N. meningitidis isolates

| Antimicrobials with Disk content |  | No (49) | % |
| --- | --- | --- | --- |
| Ceftriaxone /CRO -30µg | I | - | - |
|  | R | 18 | 36.7 |
|  | S | 31 | 63.3 |
| Ampicillin(10µg), | I | - | - |
|  | R | 42 | 83.7 |
|  | S | 8 | 16.3 |
| Amoxicillin (10µg), | I | - | - |
|  | R | 43 | 87.8 |
|  | S | 6 | 12.2 |
| Trimethoprim sulfamethoxazole1/SXT 1.25/23.75µg | I | 6 | 12.2 |
|  | R | 32 | 65.3 |
|  | S | 11 | 22.5 |
| Cefepime(30µg) | I | 3 | 11.3 |
|  | R | 7 | 15.2 |
|  | S | 36 | 73.5 |
| Meropenem(30µg) | I | - | - |
|  | R | 18 | 36.7 |
|  | S | 31 | 63.3 |
| Ceftazidime(30µg) | I | 1 | 2 |
|  | R | 22 | 44.9 |
|  | S | 26 | 53.1 |

Table 8: Bivariable logistic regression analysis for association between sex and antibiotic susceptibility pattern of 49 N. meningitidis isolates

| Antibiotics |  | Sex |  | OR (95% CI) | P-value |
| --- | --- | --- | --- | --- | --- |
|  |  | Male (n=26) | Female (n=23) |  |  |
| Ceftriaxone | S | 18 | 13 | 1.73 (0.54, 5.58) | 0.36 |
|  | R | 8 | 10 |  |  |
| Ampicillin | S | 5 | 3 | 1.58 (0.33, 7.53) | 0.56 |
|  | R | 21 | 20 |  |  |
| Amoxicillin | S | 3 | 3 | 0.87 (0.15, 4.80) | 0.87 |
|  | R | 23 | 20 |  |  |
| Trimethoprim sulfamethoxazole/SXT | S | 6 | 5 | 1.08 (0.28, 4.15) | 0.91 |
|  | R | 20 | 18 |  |  |
| Cefepime | S | 20 | 19 | 0.70 (0.17, 2.88) | 0.62 |
|  | R | 6 | 4 |  |  |
| Meropenem | S | 18 | 13 | 1.73 (0.54, 5.59) | 0.36 |
|  | R | 8 | 10 |  |  |
| Ceftazidime | S | 15 | 11 | 1.49 (0.48, 4.60) | 0.49 |
|  | R | 11 | 12 |  |  |

Table 9: Bivariable logistic regression analysis for the association between age and antibiotics susceptibility pattern of 49 N. meningitidis isolates

| Antibiotics |  | Age |  |  | OR (95% CI) | P-value |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | <10 | 11-20 | ≥21 |  |  |
| Ceftriaxone | S | 5 | 16 | 10 | 1.00 2.0 (0.54, 5.59) 1.28 (0.54, 5.59) | 0.44 0.75 |
|  | R | 4 | 10 | 4 |  |  |
| Ampicillin | S | 2 | 4 | 2 | 1.00 0.64 (0.09, 4.24) 0.58 (0.06, 5.11) | 0.64 0.63 |
|  | R | 7 | 22 | 12 |  |  |
| Amoxicillin | S | 1 | 3 | 2 | 1.00 1.04 (0.09, 11.52) 1.33 (0.10, 17.27) | 0.97 0.82 |
|  | R | 8 | 23 | 12 |  |  |
| Trimethoprim-sulfamethoxazole/SXT | S | 2 | 6 | 3 | 1.00 0.95 (0.12, 7.23) 1.05 (0.17, 6.46) | 0.96 0.95 |
|  | R | 7 | 20 | 11 |  |  |
| Cefepime | S | 8 | 19 | 12 | 1.00 0.75 (0.06, 9.72) 0.34 (0.03, 3.23) | 0.82 0.34 |
|  | R | 1 | 7 | 2 |  |  |
| Meropenem | S | 5 | 16 | 10 | 1.00 2.0 (0.34, 11.54) 1.28 (0.27, 5.93) | 0.44 0.75 |
|  | R | 4 | 10 | 4 |  |  |
| Ceftazidime | S | 3 | 14 | 9 | 1.00 3.6 (0.62, 21.03) 2.3 (0.47, 11.34) | 0.15 0.23 |
|  | R | 6 | 12 | 5 |  |  |

Figure 1: Agarose gel electrophoresis of the PCR products obtained from optimization of PCR targeting the sodC gene.

Lane 1, 1kb+ DNA ladder marker; Lane 2, 7, and 12 are N. meningitidis ATCC control strain, Lane 3, 8, and 13; 4, 9, and 14; 5, 10, and 15 are N. meningitidis  islotates of serogroup X, Y, and W, respectively amplified by three different versions of sodC primers. Lane 6, 11 and 16 are negative controls.

Figure 2: Agarose gel electrophoresis of the PCR products obtained from PCR targeting the sodC gene and the ctrA gene.

Lane 1, 1kb+ DNA ladder marker; Lane -2 N. meningitidis ATCC control strain, Lane 3-10 clinical samples amplified by primers targeting the sodC,  Lane 11, negative control.; and Lane 12–20 clinical samples amplified by primers targeting the ctrA gene.